## Slide 1
Multi-Site BC Measurement Analysis
Weekly Progress Update — April 1, 2026 New: Multi-site wavelength analysis, cross-plots, distributions, AERONET filtering

## Slide 2
What Was Done This Week
1. Downloaded & processed minute-resolution aethalometer data for all 4 sites
(Beijing: 770K rows, Delhi: 561K, JPL: 1.27M, Addis: 1.56M)
2. Built QC pipeline: negative removal + IQR outlier removal + 10-min rolling median
(10-min best matches instrument smoothing — r=0.993, wins 7/11 comparisons)
3. Multi-site diurnal wavelength analysis (all 5 wavelengths, all seasons)
4. Cross-plots: HIPS BC vs FTIR EC vs Aeth for all 4 sites
5. Distribution plots: histograms + violins for all sites
6. AERONET sequential exclusion applied to Addis cross-plots
7. Comprehensive research progress document (all 10 meetings tracked)

## Slide 3
1. Cross-Plots: Addis Is the Anomalous Site
Beijing, Delhi, JPL: near-zero intercepts, slopes close to 1:1
Addis: large positive intercept (~2 ug/m3) and compressed HIPS range
Two co-occurring problems: OFFSET + COMPRESSION
Delhi has similar concentrations but works well -> anomaly is SITE-SPECIFIC, not concentration-dependent

## Slide 4
2. Distribution Comparison: HIPS Compressed at Addis
Beijing/Delhi/JPL: FTIR EC and HIPS BC have comparable distributions
Addis: HIPS BC squeezed into narrow range (~2-4 ug/m3)
FTIR EC and Aeth span much wider (~2-15 ug/m3)
This compression is unique to Addis

## Slide 5
3. Green/IR Ratio: Addis Instrument Issue Confirmed
Green/IR at Addis = 0.55 (half of IR)
All other sites: 1.05-1.15 (expected)
Blue/Red/IR collapsed together at Addis
NOT aerosol physics — instrument or data processing issue on MA350-0238
Question for Cena: known issue? Correction available?

## Slide 6
4. Cross-Site Wavelength Comparison
UV/IR ratio: Addis highest (~1.7)
Blue/IR: Addis collapsed (~1.0), others show proper separation
AAE: all sites below literature range (eBC issue)
Delta-C: Addis + Delhi show strongest biomass signatures

## Slide 7
5. Wavelength/IR Ratio Diurnal — All Sites

## Slide 8
6. Seasonal Patterns: Delhi
Post-monsoon (Oct): dramatically higher AAE and UV/IR ratio
Consistent with crop-residue burning after Oct-Nov harvest
Monsoon (Jul-Sep): lowest AAE, washout effect
Seasonal definitions from Sebastian et al. (2022)

## Slide 9
7. Seasonal Patterns: Beijing
Winter: highest AAE (~0.52 midday)
Summer: lowest AAE (~0.15)
Matches Sun et al. (2022) exactly
Proper spectral separation at all seasons (unlike Addis)

## Slide 10
8. Seasonal Patterns: Addis Ababa
Green at 0.5x persists across ALL seasons
Kiremt evening peak distinct
Dry season morning has highest UV divergence
Divergence timing differs from concentration peaks

## Slide 11
9. AERONET Filtering: Does It Improve Addis Cross-Plots?
157 filter days matched with AERONET AOD
Sequential exclusion: remove lowest 0/20/40/60% AOD days
Tests if decoupled surface-column days explain scatter
Advisor requested this analysis specifically

## Slide 12
10. Normalized Diurnal Heatmaps
Addis Green channel shows deep blue stripe (near zero)
Beijing/Delhi/JPL show expected warm-cool gradient
Confirms instrument anomaly is persistent, not transient

## Slide 13
11. AAE Diurnal Pattern — All Sites
All sites: AAE_UV_IR = 0.15-0.63
Literature expects 1.0 (fossil fuel) to 2.0 (biomass)
Likely because data is eBC (MAC-corrected), not raw b_abs
Raw ATN columns exist — could recalculate from absorption coefficients
Addis Blue/IR AAE near zero — confirms wavelength correction issue

## Slide 14
12. What We Tried That Didn't Explain the Anomaly
1. Relative humidity thresholds -> "Red herring" (sample sizes too small)
2. Temperature (T_min, T_max) -> Counterintuitive (burning peaks when WARMER)
3. Optical saturation -> No curvature in data (linear throughout)
4. Influential point removal -> Dataset too tight
5. OC/EC ratio -> No additional insight
6. Dust (AERONET coarse AOD) -> r = -0.33 (NEGATIVE) = RULES OUT dust
Higher dust -> HIPS UNDERESTIMATES vs FTIR (opposite of expected)
7. Iron/EC ratio thresholds -> Iron not causing slope discrepancies
8. Source apportionment (PMF) -> All combustion sources show same slopes
9. Combined source grouping -> Biomass vs fossil fuel fully overlap
10. All-season aggregation -> Masks real patterns
11. Wavelength stacking -> Unique to Addis = instrument issue

## Slide 15
13. Open Questions for Cena Meeting
1. Green channel at 0.55x IR: Known issue with MA350-0238?
Correction or recalibration available?
2. AAE from eBC vs b_abs: Is data MAC-corrected?
Raw ATN exists — should we recalculate?
3. HIPS anomaly root cause: Tested humidity, temperature, saturation,
dust, iron, sources — ALL negative. What else could cause it?
4. Filter loading effects: Could high BC loading at Addis (~5 ug/m3
vs ~0.5 at JPL) affect HIPS measurement differently?
5. Seasonal wavelength data: Can we get minute-resolution for
Delhi/Beijing to extend wavelength comparison?

## Slide 16
14. Next Steps
HIGH PRIORITY:
- Paper skeleton (advisor has asked multiple times)
- Contact Amina for Delhi seasonal confirmation
- Remind advisor -> schedule Cena meeting
- Queens College comparison image
MEDIUM PRIORITY:
- Data exclusion registry (formal log)
- Literature search: filter-based absorption, PSAP, MAAP
- Warren White HIPS/IMPROVE papers
- Calculate AAE from raw b_abs (ATN data)
COMPLETED THIS WEEK:
- 4 new analysis notebooks + 1 config module
- Comprehensive progress doc covering all 10 meetings
- 23 action items tracked as complete