## Slide 1
Addis FTIR calibration
Weekly update
Spectral exclusions, seasonal analogs
and the next validation design
Ahmad Jalil
10 September 2026
Notes: Follow-up to the September 3 meeting and Ann’s approved task email. Baseline-corrected spectra throughout. All new Addis readouts in this deck are exploratory.

## Slide 2
Spectral regions for analog selection
Rank IMPROVE spectra by Pearson correlation to each group’s median Addis spectrum.
Selection space | Channels used
Full AIRSpec control | 1426–3998 cm⁻¹
Primary exclusion | Remove 1800–2500 cm⁻¹
Upper-end tests | Also remove >3600 or >3500 cm⁻¹
500 unique, TOR-eligible filters per selection
All PLS fits retain the same full AIRSpec grid.
1850 cm⁻¹ lower-bound sensitivity also checked. Raw-spectra selection stays closed.
2
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/mask_channels.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/library_eligibility.csv
Available grid: 2002 channels. Only correlation selection is masked. Signed r avoids rewarding inverted spectra. The group median preserves the earlier seasonal selection definition. Physical IMPROVE FilterIds are deduplicated after ranking. Mean-across-filter correlation is retained as a separate sensitivity run.

## Slide 3
Region exclusions change the analog sets
Number of filters replaced relative to the full-spectrum selection, out of 500.
Chart: 
| CO₂ cut | + >3600 cut | + >3500 cut
Replaced filters | 22 | 61 | 70
Season pair | Shared filters / 500 | Jaccard overlap
Dry / Belg | 6 | 0.6%
Dry / Kiremt | 1 | 0.1%
Belg / Kiremt | 160 | 19.0%
Seasonal overlap uses the primary CO₂ exclusion. Jaccard = shared filters / combined unique filters.
3
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/mask_membership_changes.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/seasonal_overlap.csv

## Slide 4
A near-zero Dry intercept hides a shallow slope
All three seasonal models fall below the previous TOR R² screening threshold of 0.85.
Season | n | Addis R² | Deming slope [95% CI] | Intercept [95% CI]
Dry | 105 | 0.552 | 0.31 [0.24, 0.36] | -0.06 [-0.26, 0.18]
Belg | 61 | 0.445 | 0.53 [0.32, 0.74] | -0.36 [-1.40, 0.68]
Kiremt | 73 | 0.646 | 0.90 [0.81, 1.09] | -0.97 [-2.24, -0.37]
Site-held-out TOR R²: 0.428–0.700
TOR tests differ between cohorts, so these are not paired model comparisons.
x = HIPS Fabs/MAC, y = FTIR EC, both µg/m³. February belongs to Dry (dry_feb).
4
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/regression_metrics.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/calibration_fits.csv
MAC=10. 95% percentile intervals from 2000 month-block bootstrap draws. PLS k comes from site-grouped CV on training sites only, first-major-minimum rule. Readouts use all seasonal Addis filters that also defined spectral selection. They are exploratory, not an independent Addis holdout. TOR R² is squared Pearson correlation, Q² prediction accuracy is separately saved in calibration_fits.csv. Error ratio lambda=3.3354.

## Slide 5
PMF models remain exploratory
102 matched Addis filters support five separate source selections.
PMF dominant source | n | Addis R² | TOR R² | Slope | Intercept
Charcoal | 22 | 0.305 | 0.686 | 0.76 | -0.70
Fossil Fuel Combustion | 25 | 0.785 | 0.792 | 0.64 | -1.11
Polluted Marine | 23 | 0.106 | 0.169 | 0.06 | 0.25
Sea Salt Mixed | 21 | 0.831 | 0.285 | 0.81 | -2.24
Wood Burning | 11 | 0.554 | 0.628 | 0.55 | 0.09
Slope and intercept use Deming (intercept in µg/m³). 137 filters have no PMF match.
5
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/pmf_overlap.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/regression_metrics.csv
GF fractions normalized with the canonical ETAD helper before dominant-source assignment. Exact calendar-date join. Dominant source does not mean pure-source aerosol. Each row uses its own selected calibration and only its own PMF target group. All 95% uncertainty intervals, common-model group readouts and TOR tests are in the analysis report.

## Slide 6
Training membership and Bishoftu lot
The historical Addis calibration used site-grouped CV and a separate TOR site test.
Historical model: lowest OC/EC 440, AIRSpec, k = 8
387 filters in the fit, 53 in the TOR test.
All 10 IMPROVE filters in the meeting’s Addis / Bishoftu figure
were outside the historical 440-filter calibration cohort.
Bishoftu: all 40 filters confirmed as lot 251 in raw metadata.
HIPS values: 26 shipped + 14 reconstructed.
6
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/historical_addis_winner_membership.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/meeting_figure_top5_membership.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/old_deck_top5_membership.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/representative_top5_membership.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/bishoftu_lot_audit.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/bishoftu_primary_lot_check.csv
New representative top-five lists use each site median, so do not claim they reproduce the old medoid display. Historical training membership is reconstructed from exact original OC/EC ordering and the locked site split seed 20260717. Distinguish membership in the 440-filter cohort from the subset actually used to fit.

## Slide 7
A separate Addis selection and validation set
Proposed design: hold out whole calendar months within each season.
Season | Selection filters | Validation filters | Validation months
Dry | 76 | 29 | 5
Belg | 40 | 21 | 3
Kiremt | 50 | 23 | 4
166 filters select the analog cohorts.
73 filters are reserved for a later prediction check.
Retrospective proposal, not yet scored. Earlier Addis exploration prevents a pristine holdout claim.
7
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/proposed_addis_split.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/split_counts.csv
Seed 20260910. Approximately one-third of each season’s observed calendar-month blocks are reserved. Lock masks, aggregation, N and k-selection before scoring. Select using only selection-side spectra, then fit using IMPROVE TOR. Neither validation spectra nor HIPS values may choose a calibration. For a stronger confirmatory test collect new samples or use a previously unexamined cohort. https://scikit-learn.org/stable/common_pitfalls.html#data-leakage https://scikit-learn.org/stable/modules/cross_validation.html

## Slide 8
What the uncertainty intervals establish
The existing Deming functions returned point estimates only. Bootstrap intervals are now available.
Intervals use whole-month resampling to keep nearby filters together.
HIPS uncertainty: 0.291 µg/m³ after MAC conversion.
The FTIR error proxy gives λ = 3.34.
An interval covering slope 1 or intercept 0 means compatibility.
It does not prove agreement or acceptable practical error.
Intervals condition on the chosen calibration and fixed MAC and λ. They exclude model-selection uncertainty.
8
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/regression_metrics.csv
Median HIPS_Uncertainty / MAC read from parameter rows. sigma_y=0.531 µg/m³ is the historical held-out TOR RMSE proxy, not a measured per-filter uncertainty. It is approximate and shared across models; sensitivity to lambda belongs in interpretation. References: https://search.r-project.org/CRAN/refmans/mcr/html/mc.bootstrap.html and https://stat.ethz.ch/CRAN/web/packages/deming/refman/deming.html. No equivalence test or practical margin has been established.

## Slide 9
The reference calibration remains the benchmark
Distinct spectral analog sets have not yet produced a stronger seasonal calibration.
Dry shares only 1–6 filters with the rainy-season analog sets.
Seasonal TOR R² ranges from 0.43 to 0.70.
The historical OC/EC reference gives 0.924 on its own TOR test.
Next: agree on the Addis split and freeze the selection rule
before assessing the reserved predictions.
Cohort-specific TOR tests differ. The mean-correlation sensitivity reaches the same screening conclusion.
9
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/calibration_fits.csv
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/ann_weekly_20260910/seasonal_overlap.csv
The primary results preserve the prior median-spectrum selection. The first pass using mean-across-filter correlation is retained under output/tables/ann_weekly_20260910_mean_sensitivity. Both approaches fail the 0.85 screening floor for each of the three season-specific fits. Do not claim statistical superiority from different TOR test populations or adopt a calibration selected on already explored Addis outcomes.

## Slide 10
546 filters have both HIPS and FTIR EC
Audit of the legacy four-site dataset. Separate from the 239-filter calibration cohort.
Site | HIPS/FTIR pairs | HIPS + IR date candidate | HIPS + reported collection window
Beijing | 163 | 73 | 0
Delhi | 63 | 25 | 0
JPL | 130 | 68 | 130
Addis Ababa | 190 | 181 | 187
Total | 546 | 347 | 317
347 date-window candidates still need matching to actual sampling intervals.
Counts precede exclusions. One registered exclusion leaves 545 HIPS/FTIR diagnostic pairs.
10
Notes: The catalog retains 1,055 site/base FilterId pairs and all 44,493 original measurement rows. The three numeric columns describe separate overlapping subsets, not sequential attrition. The 317 HIPS filters with a positive-duration reported collection window comprise 187 Addis and 130 JPL filters. Only current Addis/JPL portal files were available in the inspected directory. A zero for Beijing/Delhi means no window recovered from these inputs, not proof that metadata do not exist. Same-filter HIPS/FTIR agreement is a measurement diagnostic, not independent validation of FTIR-predicted EC.
Sources:
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/matched_sample_audit/selection_summary.parquet
/Users/ahmadjalil/github/aethmodular/docs/matched-sample-audit-2026-09-10.md

## Slide 11
The saved daily windows need rebuilding
The historical “9 AM” builder created 15:00 boundaries for three sites.
Site | Saved local timestamp | Coverage limitation
Beijing | 15:00 | Input availability only
Delhi | 15:00 | Input availability only
JPL | Mostly 15:00 | One day reaches 104.0%
Addis Ababa | 09:00 | No coverage field
The source now uses local 09:00 bounds and the actual elapsed minutes.
Saved aggregates await regeneration. Filter metadata must determine the final matching intervals.
11
Notes: Beijing has 590 timestamps at 15:00 and Delhi has 289. JPL has 762 at 15:00, two at 14:00 and one at 16:00. All three saved inputs also contain duplicate datetime_local columns. Addis has 515 timestamps at 09:00. The old builder subtracted 15 hours, aggregated midnight bins and added 15 hours. The revised source uses local [start,end) calendar intervals labelled at their end. It counts distinct valid minutes per channel and allows 23/25-hour DST days. Without upstream observation flags, non-null input availability still cannot establish observed coverage. The saved pickles have not been regenerated or relabelled.
Sources:
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/matched_sample_audit/aeth_inventory.json
/Users/ahmadjalil/github/aethmodular/docs/matched-sample-audit-2026-09-10.md

## Slide 12
Collection can span several days
Portal exports restored 452 positive-duration windows and flagged 23 timing issues.
Example filter | Collection dates (local time) | Elapsed hours | Sampled hours
ETAD-0241 | 7–15 Aug 2024 | 192 | 24
USPA-0329 | 6–16 Mar 2024 | 237 | 48
ETAD-0178 | 24 Feb 2024, same start/end | 0 | 24
The average must represent the hours when the sampler actually operated.
Sampled hours do not identify active periods. ETAD-0243 also has a one-day start-date disagreement.
12
Notes: Seven Addis and sixteen JPL portal records have sampled hours inconsistent with the elapsed collection window or invalid timing. Five Addis windows span 192–216 hours with 24 sampled hours. Fifteen JPL windows span 237–238 hours with 48 sampled hours. ETAD-0163 separately reports 22.8 sampled hours over 24 hours. USPA-0243 reports 24 sampled hours over 23 hours. The zero-length ETAD-0178 record remains flagged and does not count among the 452 positive-duration windows. ETAD-0243 starts on 2024-08-24 in the portal but has SampleDate 2024-08-25 in the unified table. Hours agreeing within 0.05 h establish metadata consistency only, not an observed on/off log.
[
{
"id": "ETAD-0241",
"start_utc": "2024-08-07 06:00:00+00:00",
"end_utc": "2024-08-15 06:00:00+00:00",
"elapsed_h": 192,
"sampled_h": 24,
"status": "sampled_hours_differ_from_elapsed_window; active_periods_needed",
"source_links": [
{
"portal_source_file": "/Users/ahmadjalil/Library/CloudStorage/GoogleDrive-ahzs645@gmail.com/My Drive/University/Research/Grad/Data/Davis Data/EC-HIPS-Aeth Comparison/Data/SPARTAN portal downloads/ETAD/FilterBased_ChemSpecPM25_ETAD.csv",
"portal_source_row": 5686
}
]
},
{
"id": "USPA-0329",
"start_utc": "2024-03-06 18:00:00+00:00",
"end_utc": "2024-03-16 15:00:00+00:00",
"elapsed_h": 237,
"sampled_h": 48,
"status": "sampled_hours_differ_from_elapsed_window; active_periods_needed",
"source_links": [
{
"portal_source_file": "/Users/ahmadjalil/Library/CloudStorage/GoogleDrive-ahzs645@gmail.com/My Drive/University/Research/Grad/Data/Davis Data/EC-HIPS-Aeth Comparison/Data/SPARTAN portal downloads/USPA/FilterBased_ChemSpecPM25_USPA.csv",
"portal_source_row": 8588
}
]
},
{
"id": "ETAD-0178",
"start_utc": "2024-02-24 06:00:00+00:00",
"end_utc": "2024-02-24 06:00:00+00:00",
"elapsed_h": 0,
"sampled_h": 24,
"status": "invalid_or_nonpositive_collection_window",
"source_links": [
{
"portal_source_file": "/Users/ahmadjalil/Library/CloudStorage/GoogleDrive-ahzs645@gmail.com/My Drive/University/Research/Grad/Data/Davis Data/EC-HIPS-Aeth Comparison/Data/SPARTAN portal downloads/ETAD/FilterBased_ChemSpecPM25_ETAD.csv",
"portal_source_row": 4667
}
]
}
]
Sources:
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/matched_sample_audit/portal_filter_metadata.parquet
/Users/ahmadjalil/github/aethmodular/docs/matched-sample-audit-2026-09-10.md

## Slide 13
EC provenance and the next comparison
500 ChemSpec EC groups contain conflicting values under the same parameter name.
ChemSpec EC repeats the FTIR EC product.
Independent EC validation requires a separate reference.
The audit preserves each original value and source row.
The conflicts still need tracing to their original parameter definitions.
Next milestone: verified sampling intervals and optical conversions,
followed by withheld-period checks of any proposed adjustment.
This section audits legacy four-site inputs. The seasonal calibration results use a separate cohort.
13
Notes: The 500 conflicting ChemSpec EC groups are Beijing 156, Addis 175, Delhi 27 and JPL 142. For example, CHTS-0658 has 0.93 and 0.06 micrograms per cubic metre under the same parameter name. The audit neither averages conflicting values nor chooses the first record. Existing repo provenance work identifies ChemSpec EC concentration as FTIR EC and ChemSpec BC as HIPS Fabs/10, subject to rounding. Before an optical comparison, verify the instrument absorption/eBC conversion history, wavelength, particle size and reference conditions. Keep each physical filter in a single validation group and assess adjustments in withheld calendar blocks. Missing inputs still include Beijing/Delhi collection metadata in the inspected source set and active sampling logs for intermittent windows. This audit does not re-estimate the separate seasonal calibration shown earlier in the presentation.
Sources:
/Users/ahmadjalil/github/aethmodular/docs/open-items.md
/Users/ahmadjalil/github/aethmodular/docs/matched-sample-audit-2026-09-10.md
/Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/tables/matched_sample_audit/selection_summary.parquet

## Slide 14
Dry: full spectral range of both sets
394 IMPROVE filters in the actual fit and all 105 Addis filters. Grey = IMPROVE. Colour = Addis.
Chart: AIRSpec-corrected absorbance
Thin lines show the full minimum–maximum envelope. Thick lines show the median. All individual traces are supplied separately.
14
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/plots/ann_weekly_20260910/full_calibration_vs_addis_Dry__Oct_Feb.png
Complete 2002-channel data in the editable chart. Full-range envelope represents all actual training spectra and every seasonal Addis spectrum. The 500 selected filters include the TOR test subset, which does not enter the fit; the chart uses only actual training filters. This display summarizes the full set rather than choosing five examples. The downloadable PNG shows every individual trace. No extra smoothing or area normalization applied.

## Slide 15
Belg: full spectral range of both sets
373 IMPROVE filters in the actual fit and all 61 Addis filters. Grey = IMPROVE. Colour = Addis.
Chart: AIRSpec-corrected absorbance
Thin lines show the full minimum–maximum envelope. Thick lines show the median. All individual traces are supplied separately.
15
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/plots/ann_weekly_20260910/full_calibration_vs_addis_Belg__Mar_May.png
Complete 2002-channel data in the editable chart. Full-range envelope represents all actual training spectra and every seasonal Addis spectrum. The 500 selected filters include the TOR test subset, which does not enter the fit; the chart uses only actual training filters. This display summarizes the full set rather than choosing five examples. The downloadable PNG shows every individual trace. No extra smoothing or area normalization applied.

## Slide 16
Kiremt: full spectral range of both sets
390 IMPROVE filters in the actual fit and all 73 Addis filters. Grey = IMPROVE. Colour = Addis.
Chart: AIRSpec-corrected absorbance
Thin lines show the full minimum–maximum envelope. Thick lines show the median. All individual traces are supplied separately.
16
Notes: /Users/ahmadjalil/github/aethmodular/research/ftir_hips_chem/output/plots/ann_weekly_20260910/full_calibration_vs_addis_Kiremt__Jun_Sep.png
Complete 2002-channel data in the editable chart. Full-range envelope represents all actual training spectra and every seasonal Addis spectrum. The 500 selected filters include the TOR test subset, which does not enter the fit; the chart uses only actual training filters. This display summarizes the full set rather than choosing five examples. The downloadable PNG shows every individual trace. No extra smoothing or area normalization applied.